#### 1. Overview

1.1 The purpose of the document: This document is used by user to install LABULK 335 Tap Density Tester.

1.2 How to make record: Tick the boxes after verification and Cross it in case of failure.

#### 2. Before Installation

2.1 Environment

2.1.1 Temperature is from 5-40 Celsius degree with relative humidity less than 85% without condensation.

2.1.2 Clear and clean without smoke, dust or any mechanical vibration or electromagnetic disturbance.

2.1.3 Solid and stable workbench, Length at least 1800mm; Width at least 500mm, Height at least 500mm and less than 1000mm.

2.1.4 Power is properly grounded. In case of Power instability, UPS (universal powder supply) is a MUST.

| Environment is OK for installation | OK |
| --- | --- |

2.2 Unpacking

2.2.1 Check the outside box, no dent, no damage.

2.2.2 Check the inside carton

2.2.3 Check LABULK 335 host and accessories according to LABULK 335 Packing List.

| All the goods inside the box are unpacked and check, all OK | OK |
| --- | --- |

3 Installation

3.1 Electronic balance with capability 600g to the accuracy of 10mg (0.01g).

3.2 Oven for samples to be dried before test if necessary.

| Accessories are ready | OK |
| --- | --- |

3.3 Put LABULK 335 on workbench and no foot hangs in the air.

| The host is placed properly | OK |
| --- | --- |

3.4 Install tapping assembly.

3.4.1 Take a glass cylinder usually 100mL/250mL, put on the ring pad. Volume of the cylinder is as per standard, please ask us for suggestions

Ring pad is installed

3.4.2 Put on Aluminium washer

Aluminium washer is installed

3.4.3 Put the round pad on bottom of the cylinder

Round pad is installed

3.4.4 Screw in tight

Screw tight

3.5 Check the desired vibration assembly tapping height, both is 3mm, please ask us for height 14 quotation.

| Height is OK | OK |
| --- | --- |

3.6 Connect balance to LABULK 335 with an RS-232 cable.

| Connected. | OK |
| --- | --- |

3.7 Turn on LABULK 335 and balance check if the light and LED are OK

| The light and LED are OK. | OK |
| --- | --- |

3.8 Touch ”Check”-“Check” to see if the communication is OK.

| Communication to Controller and Balance are OK. | OK |
| --- | --- |

4. Customer acceptance

| The Installation Qualification of LABULK 0335 Tap Density Tester has accomplished; the installation and commissioning of the system is normal and the instrument will be handed over eligible for use. | Serial Number | Field Service Engineer | Customer |
| --- | --- | --- | --- |
|  |  | Company: Engineer: Date: | Company: Engineer: Date: |