#### 1. Overview

1.1 The purpose of the document: This document is used by user to operate LABULK 0335 Tap Density Tester in order to verify all functions are OK.

1.2 How to make record: Tick the boxes after verification and Cross it in case of failure.

#### 2. Before Verification

2.1 Necessary Devices: LABULK 0335 Tap Density Tester, Balance etc.

2.2 Necessary connections: LABULK 0335 host has been connected to balance.

2.2.1 Power supply connection, turn on each device to check and confirm.

2.2.2 Data cable connection, check the communication status between LABULK 335 and balance

.

| Power supply, data cable and pipeline connection are OK | OK |
| --- | --- |

3. Hardware Verification
3.1 Tapping assembly

3.1.1 Movement: put Tapping assembly in the platform1 and 2, freely moves it up and down.

| Tapping assembly moves freely | OK |
| --- | --- |

3.1.2 Weight (Only for USP/EP, other ignore): put Tapping assembly on the balance and read the weight(450 ±10 g).

| Tapping assembly weight is OK | OK |
| --- | --- |

3.2 Graduated cylinder (250 mL)

3.2.1 Reading: readable to 2 mL.

| Graduated cylinder reading is OK | OK |
| --- | --- |

3.2.2 Weight: with a mass of 220±44 g.

| Graduated cylinder weight is OK | OK |
| --- | --- |

4. Software verification

4.1 MOTOR

4.1.1 TIME. Touch TIME to input time then OK.

| TIME function is OK | OK |
| --- | --- |

4.1.2 COUNT. Touch COUNT to input count number then OK. Touch ARROW to select SOP count number then OK.

| COUNT input function is OK | OK |
| --- | --- |

| COUNT SOP function is OK | OK |
| --- | --- |

4.1.3 RPM. Touch RPM( revolutions per minute) to input number then OK. Touch ARROW to select SOP rpm then OK.

| RPM input function is OK | OK |
| --- | --- |

| RPM SOP function is OK | OK |
| --- | --- |

4.1.4 MODE. Touch MODE to shift between COUNT and TIME.

| RPM input function is OK | OK |
| --- | --- |

4.1.5 HEIGHT. Touch HEIGHT to shift between 3 mm and 14 mm.

| HEIGHT function is OK | OK |
| --- | --- |

4.2 SAMPLE

4.2.1 NO. Touch NO. to code the sample to be tested then OK.

| NO. function is OK | OK |
| --- | --- |

4.2.2 WEIGHT. Touch WEIGHT to input then OK.

| WEIGHT function is OK | OK |
| --- | --- |

4.2.3 V0. Touch V0 to input volume before tapping.

| V0 function is OK | OK |
| --- | --- |

4.2.4 DATE. Touch DATE to set current time.

| DATE function is OK | OK |
| --- | --- |

4.3 WORK

4.3.1 V0. Touch V0 to input volume before tapping.

| V0 function is OK | OK |
| --- | --- |

4.3.2 VF. Touch VF to input volume after tapping.

| VF function is OK | OK |
| --- | --- |

4.3.3 BALANCE. Touch BALANCE to obtain weight from the balance via RS-232 cable.

| BALANCE function is OK | OK |
| --- | --- |

4.3.4 T.D. Touch T.D. to show tap density.

| T.D. function is OK | OK |
| --- | --- |

4.3.5 C.I. Touch C.I. to show compressiblilty index.

| C.I. function is OK | OK |
| --- | --- |

4.3.6 H.R. Touch H.R. to show hausner ratio.

| H.R. function is OK | OK |
| --- | --- |

4.3.7 START. Touch START to start working.

| START function is OK | OK |
| --- | --- |

4.3.8 STOP. Touch STOP to stop working.

| STOP function is OK | OK |
| --- | --- |

4.3.9 PRINT. Touch PRINT to print test report.

| PRINT function is OK | OK |
| --- | --- |

5. Customer acceptance

| The performance verification and operation training of LABULK 0335 Tap Density Tester has accomplished; the installation and commissioning of the system is normal and the instrument will be handed over eligible for use. | Serial Number | Field Service Engineer | Customer |
| --- | --- | --- | --- |
|  |  | Company: Engineer: Date: | Company: Engineer: Date: |